교정용 선재의 직경 증가를 위한

전기도금법과 무전해도금법의 비교연구

김    재    남

전남대학교 대학원 치의학과

(지도교수 : 황현식)

(초록)

본 연구는 교정용 선재의 직경을 증가시키기 위한 방법으로 무전해도금법의 이용 가능성 여부를 전기도금법과의 비교를 통해 알아보고자 시행되었다.

0.016 인치 스테인레스 스틸 교정용 선재에 도금을 위한 전처리를 시행한 후, 시중에 판매되는 무전해니켈도금액(Hessonic-Gr□, 신풍금속, 한국)을 사용하여 90□C 온도에서 0.018 인치 직경이 될 때까지 도금을 시행하였다. 무전해도금 과정 중 시간에 따른 직경 증가율을 구하는 한편, 도금 후 세 지점의 직경을 계측하여 균일성을 평가하였다. 도금 금속의 정성분석을 위하여 X-선 회절분석을 시행하는 한편, 물성검사를 시행한 후 전기도금한 경우와 각각 비교 분석하여 다음과 같은 결과를 얻었다.

1. 무전해도금한 군이 전기도금을 시행한 군보다 강성, 항복강도, 극한강도 모두 높은 경향을 보였으며 강성과 극한강도에서 유의한 차이를 나타내었다(p<0.05).

2. 무전해도금한 군은 0.00461□0.00003 mm/5min (0.00092 mm/min)의 직경증가율을 보였으며, 전기도금한 군은 0.00821□0.00015 mm/min의 직경증가율을 보였다.

3. 도금 후 세 지점의 직경을 계측하여 균일성을 평가한 결과 두 가지 도금법 모두에서 균일한 양상을 보였다.

이상의 결과로 무전해도금법을 통해 직경이 증가된 선재가 전기도금법에 의해 직경이 증가된 선재보다 기존의 선재와 가까운 물성을 보임을 알 수 있었다. 그러나 이의 임상적 적용을 위해서는 도금시간을 감소 시켜야 함을 알 수 있었다.

A Comparative Study of Electroplating and Electroless plating for Diameter Increase of Orthodontic Wire

Jae-Nam Kim

Faculty of Dental Science

Graduate School, Chonnam National University, Gwangju, Korea

( Directed by Professor Hyeon-Shik Hwang )

(Abstract)

The purpose of this study was to evaluate electroless plating as a method to increase diameter of an orthodontic wire through comparison with eletroplating.

After pretreatment for plating to the stainless steel orthodontic wire of 0.016 inch, the electoless plating was done under the condition of 90□C until the diameter of the wire was increased to 0.018 inch using Hessonic-Gr□ which was sold by Shin-Pung metal. On the process of electroless plating the diameter of wire was measured every 5 minute to examine the increasing ratio of the wire□s diameter per time unit. And to examine the uniformity, the diameter of each 3 points of the electroless-plated orthodontic wire was measured. An X-ray diffraction test was performed to analyze the nature of the plated metal and physical property test was made. Through the comparison with electroplating, following results were obtained.

1.	In stiffness, yield strength, and ultimate strength of electroless plating wire group showed higher tendency than those of electroplated wire group, and in stiffness and ultimate strength there was statistically significant difference between two groups.

2.	In the electroless plating wire group, the increasing ratio of the diameter was 0.00461□0.00003 mm/5min (0.00092 mm/min). In the electroplated wire group, that was 0.00821□0.00015 mm/min.

3.	The results of uniformity test to measure the diameter of each 3 points showed a tendency of uniformity in all the two plating method.

The results of this study suggest that the electroless plating wire is more close to the ready-made wire than electroplating wire in terams of the physical property. On the other hand, the length of plating time should be solved for the clinical application of electroless plating.