치과용 금합금의 표면처리에 따른

교정용 브라켓의 전단결합강도 변화

민   지   현

전남대학교 대학원 치의학과

(지도교수 : 박병주)

(초록)

치과용 금합금에 브라켓을 부착하는 경우 자연치에 비하여 낮은 결합강도를 보이며, 잦은 브라켓 탈락이 나타나고 있는 바, 본 연구는 여러 가지 금합금 표면처리 방법이 교정용 레진 접착제와 금합금 간의 전단결합강도에 미치는 영향을 평가하여 금합금과 브라켓 간의 부착강도 증진방법을 모색하고자 시행하였다.

치과용 금합금으로 주조된 240개의 시편을 표면처리 유무 및 방법에 따라 무처리, 샌드블라스팅 단독처리, 샌드블라스팅과 주석도금 병용처리, 그리고 샌드블라스팅과 중간접착제 병용처리의 4가지 경우와, Ortho-one, Panavia 21, Superbond C&B의 3가지 레진접착제의 조합에 의해 12군으로 나누어 브라켓을 부착하였다. 시편을 증류수에 담아 37℃ 항온 수조 속에서 24시간 동안 보관한 후, 만능물성 시험기를 이용하여 전단결합강도를 측정한 후 다음과 같은 결과를 얻었다.

1. 치과용 금합금의 표면을 처리하지 않은 경우에 비하여 표면처리한 군에서 통계적으로 유의하게 높은 전단결합강도가 나타났다.

2. 샌드블라스팅 단독처리에 비하여 주석도금 병용처리를 시행한 경우 Panavia 21에서만 유의한 결합강도 증가가 나타났다.

3. 샌드블라스팅 단독처리에 비하여 중간접착제 병용처리를 시행한 경우 모든 접착제에서 유의한 결합강도 증가가 나타났다.

4. 사용된 레진접착제에 따른 전단결합강도를 비교한 결과  Superbond C&B가 가장 높고 그 다음으로 Panavia 21, Ortho-one 순으로 나타나는 양상을 보였다.

이상의 결과는 금합금 표면에서 브라켓 부착강도를 증가시키기 위해서는 레진접착제 종류에 관계없이 샌드블라스팅과 중간접착제 병용처리가 필요함을 시사하였다.

Change of Shear Bond Strength of

Orthodontic Brackets According to

Surface Treatment on Dental Gold Alloy

Ji-Hyun Min

Faculty of Dental Science

Graduate School, Chonnam National University,

Kwangju, Korea

(Directed by Associate Professor Byung-Ju Park)

(Abstract)

A bracket bonding to the dental gold alloy shows a lower bond strength than the natural teeth, and this can be a possible source of subsequent bond failure. This study aims to evaluate the effect of various gold alloy surface treatment techniques on shear bond strength between the orthodontic adhesives and the gold alloy and to find ways of increasing the bond strength. Two hundred and forty specimens made of the dental gold alloy were divided into twelve groups based on the combination of surface treatment methods(non-surface treatment, sandblasted, sandblasted plus tin-plated, and sandblasted plus intermediate adhesive) and adhesive systems (Ortho-one, Panavia 21, Superbond C&B). The specimens with brackets bonded were placed in distilled water at 37℃ for 24 hours and shear bond strength was measured by a universal testing machine.

The results were as follows:

1. All surface-treated groups showed a significantly higher shear bond strength than non-surface-treated groups.

2. The sandblasted plus tin-plated group showed a significantly higher shear bond strength than the sandblasted group only when Panavia 21 was involved.

3.  The sandblasted plus intermediate adhesive group showed a significantly higher shear bond strength than sandblasted group regardless of the type of adhesive used.

4. Of the three resin adhesive types, the Superbond C&B showed the  highest bond strength, followed by Panavia 21 and Ortho-one.

These  findings  suggest  that  a  combination of sandblasting and

intermediate resin treatment is desirable in order to enhance bracket

bond strength regardless of adhesive types.