국산 폴리카보네이트 브라켓 윙의

파절강도에 관한 연구

선  민  규

전남대학교대학원 치의학과

(지도교수 : 황 현 식)

(초록)

심미적 교정장치에 대한 요구가 증가함에 따라 폴리카보네이트 브라켓에 대한 관심이 커지고 있으나 이는 윙이 쉽게 파절되는 문제를 갖고 있다. 본 연구는 새롭게 개발된 한 국산 브라켓의 윙의 파절강도를 측정하여 임상에서 사용할 수 있는지 평가하기 위하여 시행되었다.

실험군으로 Alice 브라켓을, 대조군으로 Spirit MB 브라켓을 사용하여 상온에서 방치한 경우와 5□와 55□에서 2,000회 열변환 시킨 경우를 나눈 다음, 만능물성시험기를 이용하여 브라켓 윙의 파절강도를 측정하고 비교분석하여 다음과 같은 결과를 얻었다.

1. 상온에서 방치한 브라켓의 경우, Spirit MB 브라켓에 비하여 Alice 브라켓 윙의 파절강도가 높게 나타났으며 통계적으로 유의한 차이를 보였다 (p < 0.01).

2. 열변환 시킨 브라켓의 경우에도, Spirit MB 브라켓에 비하여 Alice 브라켓 윙의 파절강도가 높게 나타났으며 통계적으로 유의한 차이를 보였다 (p < 0.05).

이상의 결과는 Alice 브라켓이 윙의 파절강도 면에서 임상적으로 사용가능함을 시사하였다.

Fracture strength of tie wings in

a newly-developed polycarbonate bracket

Min-Kyu Sun

Department of Dental Science

Graduate School, Chonnam National University, Gwangju, Korea

( Directed by Professor Hyeon-Shik Hwang )

(Abstract)

As the demand for esthetic orthodontic appliances has increased, a concern about the polycarbonate bracket is on the increase. However, polycarbonate bracket wing is prone to be fractured. The purpose of this study was to evaluate the clinical usefulness of a newly-developed polycarbonate bracket by measuring fracture strength of the bracket tie wings.

Alice (K.M daicom, Korea) and Spirit MB (Ormco, Orange, CA) brackets were used as an experimental and control bracket respectively. Each bracket group was divided into halves. The half was untreated and the other half was treated with total 2,000 times of thermocycling between 5□ and 55□. The fracture strength of the wing was measured by universal testing machine. The statistical analysis reve