구강스캐너를 이용한 디지털 인상 채득시

전악스캔과 부분스캔의 정확도에 대한 비교 연구

문 윤 경

전남대학교 대학원 치의학과

(지도교수 : 황현식)

(국문초록)

최근 구강스캐너를 이용한 디지털 인상 채득에 대한 관심이 증가하여 보철을 위한 부분스캔에 대한 연구는 많이 있어왔으나, 교정치료를 위하여 전악스캔으로 얻어진 영상에 대한 연구는 미미한 실정이다. 본 연구는 구강스캐너를 이용하여 채득된 전악스캔 영상과 부분스캔 영상을 비교함으로써 구강스캐너를 이용하여 채득된 전악스캔 영상의 정확도를 알아보고자 시행되었다.

제3대구치를 제외한 전 치아가 있는 20명의 성인을 대상으로 Trios 구강스캐너를 이용하여 상하악 전악스캔 영상 및 우측구치부, 우측전치부, 좌측전치부 그리고 좌측구치부 등 각 악당 4개의 부분스캔 영상을 채득한 다음, 이를 동일한 대상자에서 인상 채득을 통하여 얻은 석고모형 레이저스캔 영상과 각각 중첩하여 얻은 shell/shell deviation을 이용하여 전악스캔 영상과 부분스캔 영상의 정확도를 비교하는 한편, 전악스캔 영상의 경우 오차발생 양상 이해를 위하여 치아별로도 오차를 산출하고 분석하여 다음과 같은 결과를 얻었다.

1. 구강내 전악 또는 부분스캔으로 얻은 영상을 석고모형 레이저스캔 영상과 비교하여 shell/shell deviation을 산출한 결과, 부분스캔을 한 경우에는 0.05~0.06mm 범위의 비교적 작은 오차를 보인 반면 전악스캔을 한 경우에는 0.09mm 이상의 다소 큰 오차를 나타내었다.

2. 하악에 비하여 상악에서의 오차가 크게 나타나는 경향을 보였으며 전악스캔의 경우에는 통계적 유의차를 나타내었다.

3. 전악스캔의 경우 치아별로 오차를 산출하고 평가한 결과, 전치부에서보다 구치부에서 오차가 큰 경향을 보였다.

4. 전악스캔 영상에서 치아별 오차를 좌우간 비교한 결과, 상악에서는 좌우간 차이가 두드러지지않은 반면, 하악의 경우 구치부에서 통계적 유의차는 없었으나, 좌측에서보다 우측의 오차가 크게 나타나는 경향을 보였다.

이상의 연구 결과 구강스캐너로 치열을 전악스캔한 영상은 부분스캔한 경우에 비하여 다소 큰 오차를 보임을 그리고 이러한 오차는 구치부로 갈수록 크게 나타남을 알 수 있었다.

Accuracy of Whole Dentition Scans

Using an Intraoral Scanner;

Comparison with Segmental Scan Images

Youn-Gyeong Moon

Department of Dental Science

Graduate School of Chonnam National University

(Supervised by Professor Hyeon-Shik Hwang)

(Abstract)

Few studies are available on the accuracy of whole dentition scans using intraoral scanners whereas segmental scans are reported to be accurate in the field of prosthetic dentistry. The purpose of the present study was to assess an accuracy of whole dentition scan images obtained with a direct chairside intraoral scanning by comparing them with segmental scan images.

The intraoral scans were performed in 20 adults with no missing teeth except third molar using TRIOS scanner. Maxillary and mandibular full-arch scans were carried out while 4 segmental scans for each arch were performed to obtain right posterior, right anterior, left anterior, and left posterior segmental scan images. Whole dentition and segmental scan images were superimposed with laser scanned plaster cast image using best-fit surface feature-based registration. As the error of each intraoral scan image, shell/shell deviations were computed for whole dentition and segmental scan images, and compared between the whole dentition and each segmental scan image. In addition, shell/shell deviations were calculated also for each individual tooth in whole dentition scan images to evaluate influencing factors on the accuracy of intraoral scans.

1. The whole dentition scan images showed relatively greater errors (0.09~0.10mm) compared to segmental scan images (0.05~0.06mm).

2. The errors were greater in the maxillary scan images than in the mandibular scans.

3. The evaluation of errors according to individual tooth revealed that the errors were greater in posterior teeth than in anterior teeth.

4. In the right/left comparison of errors, right side posterior teeth showed a tendency of greater value than the left side in the lower arch scans.

The results of the present study indicate that whole dentition scanning is less accurate than segmental scanning, particularly in posterior area.